## Slide 1
1680
1930
1960
1515
Optical Microscopy

## Slide 2
Human Hair
80 µm
Escherichia coli
2 µm
Lymphocyte
7 µm
Viruses
0.02 -0.3 µm
Pollen
40 µm
Micro ---- "small"
Skopeîn --- "see"
Optical Microscopy

## Slide 3
Optical Microscopy
Brightfield
Phase contrast
DIC
Fluorescent
Absorption
Fluorescence
Polarization
Refraction
Phase shift

## Slide 4
O
F
F
A
B
B1
A1
Magnifying

## Slide 5
Magnifying

## Slide 6
Magnifying

## Slide 7
Lens

## Slide 8
  Chromatic aberrations arise from variations in the refractive indices of the wide range of frequencies in visible light
Chromatic aberrations

## Slide 9
Chromatic aberrations
Convex
Concave

## Slide 10
Aberrations are related to the spherical nature of the lens
Geometrical or Spherical aberrations

## Slide 11
Objectives
Leeuwenhoek 1680
Zeiss 2000
Lens
Triplet
Group
Lens
Triplet
Group
Lens
Doublet
Groups
Miniscus
Lens
Hemispherical
Front Lens
Planapochromat
$uper corrected

## Slide 12
Optical aberrations

## Slide 13
Immersion Oil
c - the speed of light in vacuum
v - the speed of light in the substance
Refractive Index n =
C
v
Glass n = 1.52
Air n = 1.000
Immersion Oil
n = 1.515

## Slide 14
Lighting

## Slide 15
Lighting
critical to quantitative microscopy
Neutral Density filter
Daylight balancing filter
CONDENSER

## Slide 16
Aplanatic-Achromatic
CONDENSER
Lighting
Condencer
Top
Lens
Internal
Lens
Elements
Aperture
Diaphragm

## Slide 17
Microscope
Eyepiece
Barrel
Focusing Screw
Objective
Hooke 1670
Olympus BX 2000
Cooling
Vents
Episcopic
Lamphouse
Lamp
Collector
Lens
Microscope
Control
Circuit
Board
Diascopic
Lamphouse
Line
Cord
Transformer
Collector
Lens
Focus
Knob
Filters
Mirror
Illumination
Intensity
Adjustement
Base
Field
Diaphragm
Condenser
Mecanical
Stage
Objective
Breathshield
Vertical
Illuminator
Dichroic
Filters
Beamsplitters
Eyepiece
(Ocular)
Rubber
Eyecup
Lenses
Nose-
piece
Extension
Tube
Shutter
Focusing
Telescope
Camera
Black Box
Film Plan
Electrical Contacts
Automatic
Exposure
Monitor
Light
Sensor
Projection
Eyepiece
Trinocular
Head
Prisms
Shutter

## Slide 18
Color Fidelity

## Slide 19
Color Fidelity
CCD Detector

## Slide 20
Color Fidelity
Monitor Calibration
Calibration Slide

## Slide 21
Brightfield
Mucicarmine stain

## Slide 22
Optical Microscopy
Brightfield
Phase contrast
DIC
Fluorescent
Absorption
Fluorescence
Polarization
Refraction
Phase shift

## Slide 23
400 nm
Fluorescence
Jablonski
diagram
450 nm

## Slide 24
Incoming
photon
K
L
M
N
Fluorescence
Fluorescence

## Slide 25
Light-Tissue interaction
RAMAN
AutoFluorescence
Absorbance
Fluorescence
Stain
Fixation (or not)

## Slide 26
Fluorescence
Mucicarmine stain

## Slide 27
Fluorescence
Mucicarmine stain

## Slide 28
Stain -Tissue interaction
Tissue
+
_
Stain
Ionic
bonds
+
_
Hydrogen
bonds
Van der Waals
forces
Covalent
bonds
electron delocalization
FRET
Fluorescence-lifetime
changes
Fluorescence spectra modulations
Quenching
characterize the molecular microenvironment adjacent to ...

## Slide 29
Fluorescence
Mucicarmine stain
Local Discriminant Bases
Nearest-Neighbor classification
Spectral classifier
Hierarchical fusion using vector quantization
Unmixing-based constrained nonnegative matrix factorization
Spectral Weighting Envelopes
1BT Based Band Selection
Spectral Angle Mapper
Complete Linear Spectral Unmixing
Minimum Noise Fraction Transformation
Principal Component Transformation

## Slide 30
Hyperspectral
Notes: Journal of Biomedical Optics 19(1), 010901 (January 2014)
Medical hyperspectral imaging: a review

## Slide 31
Specimen Thickness

## Slide 32
Point Spread Function
Z
Working distance

## Slide 33
Point Spread Function
PSF describes the imaging system response to a point input
Z
In microscopy the point spread functions is asymmetric due to lens imperfections

## Slide 34
Deconvolution
Optical sectioning

## Slide 35
Manz at al. J Microbiol Methods 2000, 40:125-34.
Fig. * Tissue sections of Chondrosia reniformis after hybridization with the Bacteria specific probe EUB338
AutoDeblur from Media Cybernetics
Deconvolution

## Slide 36
Focus

## Slide 38
Dr. Marvin Minsky
Confocal microscope: 1957
Confocal microscope

## Slide 39
LASER
Light
Amplification by the
Stimulated
Emission of
Radiation
Scanning
Confocal
Microscope
Laser
Confocal microscope

## Slide 40
Dichroic
Mirror
Confocal microscope

## Slide 41
Conventional / Confocal

## Slide 42
Conventional / Confocal

## Slide 43
Light Sheet Microscopy
Single Plane Illumination Microscopy (SPIM)
Medaka juveniles
Philipp Keller, Stelzer Group, EMBL
Confocal microscope

## Slide 44
Z Compensate System
Confocal Olympus FV1000
543 nm
633 nm
453 nm
488 nm
515 nm
Z
Y
X
l
Time
Fluorescence
- Lifetime
Hexadimensional
(space- time)
Microscopy
Hooke 1670

## Slide 45
Confocal PSF
WF
CF

## Slide 46
400 nm
Fluorescence
Jablonski
diagram
450 nm

## Slide 47
Wavelength and Light Properties
575 nm
E = h c /l
(l)
High
Energy
Low
Energy
0,5* *2

## Slide 48
Multiphoton exitation
800 nm
800 nm
1200 nm
1200 nm
1200 nm
400 nm
Dr. Maria Göppert-Mayer : theory of two-photon quantum transitions
(two-photon absorption and emission) 1931,
Jablonski
diagram
450 nm

## Slide 49
LASER
500 nm
Pulsed LASER
1000 nm
1200 nm
800 nm
nm
500
nm
800
1200
Temporal photon concentration
Wavelength Tunability

## Slide 50
Prof. Watt W. Webb et al.
Two-photon laser scanning
fluorescence microscopy: 1990
Spatial photon concentration

## Slide 51
Confocal / Multiphoton

## Slide 52
Luo at al. Cell Structure and Function 2006, 31: 63
Comparison of photoactivation of PA-GFP in vivo with single-photon (405 nm) and multiphoton (790 nm) laser light.
Photoconversion
Excitation area

## Slide 53
Thorsten at al. Nature 2004; 427: 159
T-cell priming by dendritic cells in lymph nodes occurs in three distinct phases
In VIVO Imaging

## Slide 54
Conventional / Confocal / Biphoton

## Slide 55
Anisotropic optical properties of molecules
Multiphoton polarization microscopy
polarizator
“Biphotonic”
Laser
Linear dichroism

## Slide 56
Biphoton polarization microscopy
Cell expressing GAP43-CFP-Gαi2 and α2a-adrenergic receptor
G-proteins orientation
+ Norepinephrine
Base line
Lazar J et al. Nat Methods. 2013
Two-photon polarization microscopy reveals protein structure and function

## Slide 57
O
F
F
A
B
B1
A1
Magnifying

## Slide 58
)))))))))))))
)))))))))))))
))))))))
)))))))))))))
Diffraction

## Slide 59
Airy disk
Diffraction

## Slide 60
Resolution
?
?

## Slide 61
Resolution
Abbe diffraction limit
Abbe Resolutionx,y = λ/2NA
Numerical Aperture
NA = n•sin(θ)
n - refractive index of the imaging medium ( air, oil)
θ - aperture angle (1,4 in the best case)
D = 0.2 µm

## Slide 62
YFP
YFP
CFP
CFP
FRET
1-10 nm
Förster Resonance Energy Transfer (FRET)
Molecular Basis

## Slide 63
Z Compensate System
Confocal Olympus FV1000
543 nm
633 nm
453 nm
488 nm
515 nm
Z
Y
X
l
Time
Fluorescence
- Lifetime
Hexadimensional
(space- time)
Microscopy
Hooke 1670

## Slide 64
400 nm
Fluorescence
Jablonski
diagram
450 nm

## Slide 65
400 nm
Fluorescence (true)
Jablonski
diagram
450 nm

## Slide 66
FL1
F
FL2
FL1
FRET
1-10 nm
Time (ns)
Fluorescence-Lifetime Imaging (FLIM)
Free
Coupled

## Slide 67
Fluorescence-Lifetime Imaging (FLIM)
Lin HJ et al. Cytometry A. 2003
Fluorescence lifetime-resolved pH imaging of living cells.
Alpy F et al. J Cell Sci. 2013
STARD3 or STARD3NL and VAP form a novel molecular tether between late endosomes and the ER.
Molecular Interactions
Intracellular pH

## Slide 68
Near-field optical microscopy

## Slide 69
Near-field optical microscopy
special ($)1.78 refractive index coverslips
special ($ $) 1.78 refractive index oil
special ($ $ $) objective 100x 1.65NA

## Slide 70
Near-field optical microscopy

## Slide 71
Total internal reflection
Θ1
Θ2
The critical angle is the angle of incidence above which the total internal reflectance occurs
Evanescent wave
≈100 nm

## Slide 72
Total Internal Reflection Fluorescence Microscopy (TIRF)
McKinney et al.Nature Methods 6, 131 - 133 (2009)
5 µm
1 µm
Up to 20 nm of lateral resolution and 2–5 nm of axial resolution

## Slide 73
Structured Illumination Microscopy
The Moire effect
+
=
-
=

## Slide 74
Structured Illumination Microscopy

## Slide 75
Structured Illumination Microscopy
Up to 100 nm of lateral resolution and 300 nm of axial resolution

## Slide 76
Switch to nonfluorescente state
400 nm
Triplet state
Non fluorescent
Fluorescence

## Slide 77
PA-GFP
Emission
Desactivation
Emission
Desactivation
Super-resolution
Optical Fluctuation Imaging (SOFI)

## Slide 78
Super-resolution
Optical Fluctuation Imaging (SOFI)
Fast, background-free, 3D super-resolution optical fluctuation imaging (SOFI).
Dertinger T, Colyer R, Iyer G, Weiss S, Enderlein J.
Proc Natl Acad Sci U S A. 2009
The second-order correlation function
t1
:
:
:
:
tn
t1
:
:
:
:
tn

## Slide 79
Super-resolution
Optical Fluctuation Imaging (SOFI)
Fast, background-free, 3D super-resolution optical fluctuation imaging (SOFI).
Dertinger T, Colyer R, Iyer G, Weiss S, Enderlein J.
Proc Natl Acad Sci U S A. 2009
Up to 50 nm of lateral resolution and ? nm of axial resolution

## Slide 80
Super-resolution
Bleaching Assisted Localization Microscopy(BALM)
t1
:
:
:
:
tn
t1
:
:
:
:
tn
Fast, Bleaching/blinking assisted localization microscopy for
superresolution imaging using standard fluorescent molecules.
Burnette DT, Sengupta P, Dai Y, Lippincott-Schwartz J, Kachar B.
Proc Natl Acad Sci U S A. 2011 Dec 27;108(52):21081-6

## Slide 81
Up to 50 nm of lateral resolution and ? nm of axial resolution
Super-resolution
Bleaching Assisted Localization Microscopy(BALM)
Fast, Bleaching/blinking assisted localization microscopy for
superresolution imaging using standard fluorescent molecules.
Burnette DT, Sengupta P, Dai Y, Lippincott-Schwartz J, Kachar B.
Proc Natl Acad Sci U S A. 2011 Dec 27;108(52):21081-6

## Slide 82
doughnut-shape
+
=
Stimulated emission depletion
488 nm
520 nm

## Slide 83
Stimulated emission depletion
(STED)
Up to 50 nm of lateral resolution and 500 nm of axial resolution

## Slide 84
+
=
X Y
X Y Z
Stimulated emission depletion 3X

## Slide 85
Stimulated emission depletion
(STED 3X)
Up to 50 nm of lateral and axial resolution

## Slide 86
Fluorescence Localization Microscopy

## Slide 87
PA-GFP
PA-GFP
Fluorescence Photoactivation Localization
Microscopy
PA-GFP
Emission
Localization
(calculation)
Photobleaching
Activation

## Slide 88
Fluorescence Localization Microscopy

## Slide 89
Fluorescence Photoactivation Localization
Microscopy (PALM)
Hess, S.T., T.P. Girirajan, and M.D. Mason. 2006. Ultra-high resolution imaging by fluorescence photoactivation localization microscopy. Biophys J. 91(11):

## Slide 90
ZHUANG LAB/HARVARD UNIV.
Fluorescence Photoactivation Localization
Microscopy
Up to 30 nm of lateral resolution and 150 nm of axial resolution

## Slide 91
Switch to nonfluorescente state
550 nm
Triplet state
Non fluorescent
Fluorescence
Jablonski
diagram
600 nm
400 nm
Thiols
(R-SH)
Ground state depletion
Ground state

## Slide 92
FITC
FITC
Ground State Depletion Microscopy
direct Stochastic Optical Reconstruction Microscopy
FITC
Stochastic Activation
Localization
(calculation)
Emission
Total Deactivation
(Ground State Depletion)

## Slide 93
Ground State Depletion Microscopy

## Slide 94
Biplan Localization Microscopy
+ 400 nm
- 400 nm
0
Cylindrical
lens

## Slide 95
Biplan Localization Microscopy
Vutara, Inc..

## Slide 96
Schermelleh L et al. J Cell Biol 2010;190:165-175
Super-Resolution Microscopy
50
100
gSTED 3X
___
_____
2013
___
50
___
Biplan

## Slide 97
Merci !
Merci !

## Slide 98
http://zeiss-campus.magnet.fsu.edu/index.html
http://www.research.utah.edu/advanced-microscopy/education/index.html
http://www.olympusmicro.com/primer/anatomy/anatomy.html
http://www.microscopyu.com/articles/formulas/